## SNCImport
CBCT - Gantry Iso X (mm) | -0.33
CBCT - Gantry Iso Y (mm) | 0.065
CBCT - Gantry Iso Z (mm) | -0.06625
Table - Gantry Iso X (mm) | 0.016226
Table - Gantry Iso Z (mm) | 0.2686
Gantry Flex (mm) | 0.49
Col-Gantry misalignment (mm) | 0.0575
MLC offset (mm) | -0.2525
Table Isocentricity (mm) | 0.2143
Gantry Isocentricity (mm) | 0.25933
Collimator Isocentricity (mm) | 0.15305

## Data
Winston-Lutz Field Data | Data Date: 2024-10-15 17:47:43 | Analysis Date: 2024-10-15 17:55:17
machine id | field name | status | table yaw | gantry angle | coll angle | X offset corrected box-ball | Y offset corrected box-ball | XY offset corrected | X box center corrected | Y box center corrected | X tongue and groove correction | Y tongue and groove correction | X ball center | Y ball center | box left uncorrected | box right uncorrected | box top uncorrected | box bottom uncorrected | X box center uncorrected | Y box center uncorrected | Patient ID | Patient Name | Instance (slice) UID | Series UID
TB5 | G180 C000 0:00 | Passed | 359.959 | 179.9967 | 359.9998 | -0.18769 | 0.10911 | 0.21709 | 15.5129 | 15.2172 | 0 | 0 | 15.7006 | 15.1081 | 5.5079 | 25.5179 | 5.5569 | 24.8775 | 15.5129 | 15.2172 | QASRSWL2 | $QASRS-TX5^Winston-Lutz | 1.2.246.352.62.1.5501163403191641537.5185906158874397588 | 1.2.246.352.62.2.5591855550814472668.16651016134416038585
TB5 | G180 C090 0:21 | Passed | 359.959 | 179.9967 | 89.9999 | -0.3503 | 0.049798 | 0.35382 | 15.1468 | 15.6224 | 0 | 0 | 15.4971 | 15.5726 | 5.4915 | 24.8022 | 5.6151 | 25.6297 | 15.1468 | 15.6224 | QASRSWL2 | $QASRS-TX5^Winston-Lutz | 1.2.246.352.62.1.5022932505202810736.18176117736826310311 | 1.2.246.352.62.2.5591855550814472668.16651016134416038585
TB5 | G090 C090 0:46 | Passed | 359.959 | 90.0614 | 90 | -0.095365 | -0.19835 | 0.22008 | 15.1679 | 15.6826 | 0 | 0 | 15.2633 | 15.881 | 5.5465 | 24.7893 | 5.6763 | 25.689 | 15.1679 | 15.6826 | QASRSWL2 | $QASRS-TX5^Winston-Lutz | 1.2.246.352.62.1.5215168229937650413.11559825400615409804 | 1.2.246.352.62.2.5591855550814472668.16651016134416038585
TB5 | G000 C090 1:11 | Passed | 359.959 | 359.9602 | 89.9999 | 0.38259 | -0.43707 | 0.58086 | 15.215 | 15.5961 | 0 | 0 | 14.8324 | 16.0332 | 5.5925 | 24.8375 | 5.5836 | 25.6087 | 15.215 | 15.5961 | QASRSWL2 | $QASRS-TX5^Winston-Lutz | 1.2.246.352.62.1.5122434876230622122.12908282298789640859 | 1.2.246.352.62.2.5591855550814472668.16651016134416038585
TB5 | G270 C090 1:35 | Passed | 359.959 | 269.9386 | 89.9998 | 0.025618 | -0.17735 | 0.17919 | 15.1155 | 15.557 | 0 | 0 | 15.0899 | 15.7344 | 5.4859 | 24.7452 | 5.56 | 25.554 | 15.1155 | 15.557 | QASRSWL2 | $QASRS-TX5^Winston-Lutz | 1.2.246.352.62.1.5151400910807266628.984014451812874905 | 1.2.246.352.62.2.5591855550814472668.16651016134416038585
TB5 | G270 C270 1:56 | Passed | 359.959 | 269.9917 | 270.0001 | 0.17824 | 0.10489 | 0.20681 | 15.2755 | 15.6034 | 0 | 0 | 15.0972 | 15.4985 | 5.6414 | 24.9095 | 5.6131 | 25.5937 | 15.2755 | 15.6034 | QASRSWL2 | $QASRS-TX5^Winston-Lutz | 1.2.246.352.62.1.5594146986960155985.3615940627330091690 | 1.2.246.352.62.2.5591855550814472668.16651016134416038585
TB5 | G000 C270 2:28 | Passed | 359.959 | 360 | 270.0001 | 0.43473 | -0.19504 | 0.47648 | 15.2831 | 15.6263 | 0 | 0 | 14.8484 | 15.8213 | 5.6499 | 24.9163 | 5.6285 | 25.6241 | 15.2831 | 15.6263 | QASRSWL2 | $QASRS-TX5^Winston-Lutz | 1.2.246.352.62.1.4653000383632170084.12792974839094954936 | 1.2.246.352.62.2.5591855550814472668.16651016134416038585
TB5 | G090 C270 2:58 | Passed | 359.959 | 90.0349 | 270.0001 | 0.047339 | 0.051363 | 0.069851 | 15.0784 | 15.4462 | 0 | 0 | 15.031 | 15.3949 | 5.4465 | 24.7103 | 5.4658 | 25.4267 | 15.0784 | 15.4462 | QASRSWL2 | $QASRS-TX5^Winston-Lutz | 1.2.246.352.62.1.5466850920194951903.16537438642280104085 | 1.2.246.352.62.2.5591855550814472668.16651016134416038585
TB5 | G180 C270 3:24 | Passed | 359.959 | 180.0033 | 270.0001 | -0.16098 | 0.2933 | 0.33457 | 15.0736 | 15.6171 | 0 | 0 | 15.2346 | 15.3238 | 5.4459 | 24.7013 | 5.6323 | 25.602 | 15.0736 | 15.6171 | QASRSWL2 | $QASRS-TX5^Winston-Lutz | 1.2.246.352.62.1.5479858720157641572.17315255651440607106 | 1.2.246.352.62.2.5591855550814472668.16651016134416038585
TB5 | G180 C270 4:15 | Passed | 89.957 | 180.0033 | 270.0001 | 0.46813 | 0.21746 | 0.51617 | 15.0712 | 15.6146 | 0 | 0 | 14.6031 | 15.3971 | 5.4443 | 24.6981 | 5.6317 | 25.5975 | 15.0712 | 15.6146 | QASRSWL2 | $QASRS-TX5^Winston-Lutz | 1.2.246.352.62.1.4796884726310680617.10944648517499340976 | 1.2.246.352.62.2.5591855550814472668.16651016134416038585
TB5 | G180 C270 4:46 | Passed | 59.959 | 180.0033 | 270.0001 | 0.51956 | 0.34932 | 0.62607 | 15.0654 | 15.6255 | 0 | 0 | 14.5458 | 15.2762 | 5.4319 | 24.6988 | 5.643 | 25.608 | 15.0654 | 15.6255 | QASRSWL2 | $QASRS-TX5^Winston-Lutz | 1.2.246.352.62.1.5567567661898746564.1679050013640922288 | 1.2.246.352.62.2.5591855550814472668.16651016134416038585
TB5 | G180 C270 5:07 | Passed | 29.9531 | 180.0033 | 270.0001 | 0.34582 | 0.35825 | 0.49793 | 15.0679 | 15.6313 | 0 | 0 | 14.722 | 15.2731 | 5.4387 | 24.697 | 5.6415 | 25.6212 | 15.0679 | 15.6313 | QASRSWL2 | $QASRS-TX5^Winston-Lutz | 1.2.246.352.62.1.5132644864718072394.9831637297503514016 | 1.2.246.352.62.2.5591855550814472668.16651016134416038585
TB5 | G180 C270 5:41 | Passed | 329.959 | 180.0033 | 270.0001 | -0.3586 | 0.47259 | 0.59325 | 15.078 | 15.6314 | 0 | 0 | 15.4366 | 15.1588 | 5.4472 | 24.7087 | 5.6397 | 25.6231 | 15.078 | 15.6314 | QASRSWL2 | $QASRS-TX5^Winston-Lutz | 1.2.246.352.62.1.5466391627421975314.2943708566449341357 | 1.2.246.352.62.2.5591855550814472668.16651016134416038585
TB5 | G180 C270 6:01 | Passed | 299.959 | 180.0033 | 270.0001 | -0.24864 | 0.56975 | 0.62164 | 15.0771 | 15.6237 | 0 | 0 | 15.3258 | 15.054 | 5.4452 | 24.7091 | 5.638 | 25.6095 | 15.0771 | 15.6237 | QASRSWL2 | $QASRS-TX5^Winston-Lutz | 1.2.246.352.62.1.5138739043602403820.4139278402537124792 | 1.2.246.352.62.2.5591855550814472668.16651016134416038585
TB5 | G180 C270 6:26 | Passed | 269.959 | 180.0033 | 270.0001 | -0.36301 | 0.74013 | 0.82435 | 15.2952 | 15.6155 | 0 | 0 | 15.6583 | 14.8754 | 5.669 | 24.9215 | 5.6311 | 25.6 | 15.2952 | 15.6155 | QASRSWL2 | $QASRS-TX5^Winston-Lutz | 1.2.246.352.62.1.5744035648799867708.5834260051927852479 | 1.2.246.352.62.2.5591855550814472668.16651016134416038585
ballCBCT
x (mm) | y (mm) | z (mm)
0 | 0 | 0
0 | 0 | 0 | AcqIsocenter
0 | 0 | 0 | Ball

## Preprocess
Winston-Lutz Field Data | Data Date: 2024-10-15 17:47:43 | Analysis Date: 2024-10-15 17:55:17
machine id | field name | status | gantry angle | coll angle | table yaw | X offset corrected box-ball | Y offset corrected box-ball | XY offset corrected | X box center corrected | Y box center corrected | X tongue and groove correction | Y tongue and groove correction | X ball center | Y ball center | box left uncorrected | box right uncorrected | box top uncorrected | box bottom uncorrected | X box center uncorrected | Y box center uncorrected | Patient ID
TB5 | G180 C000 0:00 | Passed | 180 | 0 | 0 | -0.19 | 0.11 | 0.22 | 15.51 | 15.22 | 0 | 0 | 15.7 | 15.11 | 5.51 | 25.52 | 5.56 | 24.88 | 15.51 | 15.22 | QASRSWL2
TB5 | G180 C090 0:21 | Passed | 180 | 90 | 0 | -0.35 | 0.05 | 0.35 | 15.15 | 15.62 | 0 | 0 | 15.5 | 15.57 | 5.49 | 24.8 | 5.62 | 25.63 | 15.15 | 15.62 | QASRSWL2
TB5 | G090 C090 0:46 | Passed | 90 | 90 | 0 | -0.1 | -0.2 | 0.22 | 15.17 | 15.68 | 0 | 0 | 15.26 | 15.88 | 5.55 | 24.79 | 5.68 | 25.69 | 15.17 | 15.68 | QASRSWL2
TB5 | G000 C090 1:11 | Passed | 0 | 90 | 0 | 0.38 | -0.44 | 0.58 | 15.22 | 15.6 | 0 | 0 | 14.83 | 16.03 | 5.59 | 24.84 | 5.58 | 25.61 | 15.22 | 15.6 | QASRSWL2
TB5 | G270 C090 1:35 | Passed | 270 | 90 | 0 | 0.03 | -0.18 | 0.18 | 15.12 | 15.56 | 0 | 0 | 15.09 | 15.73 | 5.49 | 24.75 | 5.56 | 25.55 | 15.12 | 15.56 | QASRSWL2
TB5 | G270 C270 1:56 | Passed | 270 | 270 | 0 | 0.18 | 0.1 | 0.21 | 15.28 | 15.6 | 0 | 0 | 15.1 | 15.5 | 5.64 | 24.91 | 5.61 | 25.59 | 15.28 | 15.6 | QASRSWL2
TB5 | G000 C270 2:28 | Passed | 0 | 270 | 0 | 0.43 | -0.2 | 0.48 | 15.28 | 15.63 | 0 | 0 | 14.85 | 15.82 | 5.65 | 24.92 | 5.63 | 25.62 | 15.28 | 15.63 | QASRSWL2
TB5 | G090 C270 2:58 | Passed | 90 | 270 | 0 | 0.05 | 0.05 | 0.07 | 15.08 | 15.45 | 0 | 0 | 15.03 | 15.39 | 5.45 | 24.71 | 5.47 | 25.43 | 15.08 | 15.45 | QASRSWL2
TB5 | G180 C270 3:24 | Passed | 180 | 270 | 0 | -0.16 | 0.29 | 0.33 | 15.07 | 15.62 | 0 | 0 | 15.23 | 15.32 | 5.45 | 24.7 | 5.63 | 25.6 | 15.07 | 15.62 | QASRSWL2
TB5 | G180 C270 4:15 | Passed | 180 | 270 | 270 | 0.47 | 0.22 | 0.52 | 15.07 | 15.61 | 0 | 0 | 14.6 | 15.4 | 5.44 | 24.7 | 5.63 | 25.6 | 15.07 | 15.61 | QASRSWL2
TB5 | G180 C270 4:46 | Passed | 180 | 270 | 300 | 0.52 | 0.35 | 0.63 | 15.07 | 15.63 | 0 | 0 | 14.55 | 15.28 | 5.43 | 24.7 | 5.64 | 25.61 | 15.07 | 15.63 | QASRSWL2
TB5 | G180 C270 5:07 | Passed | 180 | 270 | 330 | 0.35 | 0.36 | 0.5 | 15.07 | 15.63 | 0 | 0 | 14.72 | 15.27 | 5.44 | 24.7 | 5.64 | 25.62 | 15.07 | 15.63 | QASRSWL2
TB5 | G180 C270 5:41 | Passed | 180 | 270 | 30 | -0.36 | 0.47 | 0.59 | 15.08 | 15.63 | 0 | 0 | 15.44 | 15.16 | 5.45 | 24.71 | 5.64 | 25.62 | 15.08 | 15.63 | QASRSWL2
TB5 | G180 C270 6:01 | Passed | 180 | 270 | 60 | -0.25 | 0.57 | 0.62 | 15.08 | 15.62 | 0 | 0 | 15.33 | 15.05 | 5.45 | 24.71 | 5.64 | 25.61 | 15.08 | 15.62 | QASRSWL2
TB5 | G180 C270 6:26 | Passed | 180 | 270 | 90 | -0.36 | 0.74 | 0.82 | 15.3 | 15.62 | 0 | 0 | 15.66 | 14.88 | 5.67 | 24.92 | 5.63 | 25.6 | 15.3 | 15.62 | QASRSWL2
Sorted |  |  | gantry angle | coll angle | table yaw | X offset corrected box-ball | Y offset corrected box-ball
|  |  | 0 | 90 | 0 | 0.38 | -0.44
|  |  | 0 | 270 | 0 | 0.43 | -0.2
|  |  | 90 | 90 | 0 | -0.1 | -0.2
|  |  | 90 | 270 | 0 | 0.05 | 0.05
|  |  | 180 | 0 | 0 | -0.19 | 0.11
|  |  | 180 | 90 | 0 | -0.35 | 0.05
|  |  | 180 | 270 | 0 | -0.16 | 0.29
|  |  | 270 | 90 | 0 | 0.03 | -0.18
|  |  | 270 | 270 | 0 | 0.18 | 0.1
|  |  | 180 | 270 | 30 | -0.36 | 0.47
|  |  | 180 | 270 | 60 | -0.25 | 0.57
|  |  | 180 | 270 | 90 | -0.36 | 0.74
|  |  | 180 | 270 | 270 | 0.47 | 0.22
|  |  | 180 | 270 | 300 | 0.52 | 0.35
|  |  | 180 | 270 | 330 | 0.35 | 0.36

## Analysis
Gantry | [17 empty] | MLC offset | -0.2525
Gantry angle | Collimator angle | Table angle | X offset corrected box-ball | Y offset corrected box-ball | CA-X | CA-Y | CA-Z | Coll-X | Coll-Y | Coll-Z | ISO-X | ISO-Y | ISO-Z | Gantry flex | Coll-Gantry-misalign | MLC-dx | MLC-dy | -0.135 | -0.2525 |  | X (mm) | Y (mm) | Z (mm)
0 | 90 | 0 | 0.38 | -0.44 | 0.38 |  | 0.44 | 0.405 |  | 0.32 | 0.33 | -0.065 | 0.06625 | 0.49 | 0.0575 | -0.025 | -0.12 | -0.0675 | -0.12625 | CBCT origin relative to BB at table zero | 0 | 0 | 0
0 | 270 | 0 | 0.43 | -0.2 | 0.43 |  | 0.2 | [8 empty] | 0.025 | 0.12 | 0.0675 | 0.12625 | Gantry isocenter relative to BB at table zero | 0.33 | -0.065 | 0.06625
90 | 90 | 0 | -0.1 | -0.2 |  | -0.1 | 0.2 |  | -0.025 | 0.075 | [5 empty] | -0.075 | -0.125 |  |  | Table axis relative to BB at table zero | 0.34623 |  | 0.33485
90 | 270 | 0 | 0.05 | 0.05 |  | 0.05 | -0.05 |  |  |  |  | ΔX | ΔY | ΔZ |  | 0.075 | 0.125
180 | 0 | 0 | -0.19 | 0.11 | 0.19 |  | -0.11 |  |  |  |  | 0.15 | 0.08 | 0.49 |  | 0.065 | -0.06
180 | 90 | 0 | -0.35 | 0.05 | 0.35 |  | -0.05 | 0.255 |  | -0.17 |  |  |  | 0.25933 | Gantry Isocentricity | -0.095 | -0.12
180 | 270 | 0 | -0.16 | 0.29 | 0.16 |  | -0.29 | [8 empty] | 0.095 | 0.12
270 | 90 | 0 | 0.03 | -0.18 |  | -0.03 | 0.18 |  | -0.105 | 0.04 | [5 empty] | -0.075 | -0.14
270 | 270 | 0 | 0.18 | 0.1 |  | -0.18 | -0.1 | [8 empty] | 0.075 | 0.14
[10 empty] | 0.32736 |  |  |  | Max square of BB displacement |  | -> | 0.045926 | Solve for smallest max (Couch Isocentricity)
Table |  |  |  |  | CA-X | CA-Z | BB-X | BB-Z | BB-X' | BB-Z' | dX | dZ | Table-X | Table-Z | BB-X" | BB-Z" | BB-R"^2
180 | 270 | 0 | -0.16 | 0.29 | 0.16 | -0.29 | 0 | 0 | 0.20125 | 0.35064 | 0.20125 | 0.35064 | 0.34623 | 0.33485 | -0.14498 | 0.015788 | 0.021268
180 | 270 | 30 | -0.36 | 0.47 | 0.36 | -0.47 | -0.2 | 0.18 | 0.1496 | 0.38304 | 0.3496 | 0.20304 |  |  | -0.19662 | 0.048187 | 0.040982
180 | 270 | 60 | -0.25 | 0.57 | 0.25 | -0.57 | -0.09 | 0.28 | 0.31429 | 0.28103 | 0.40429 | 0.001035 |  |  | -0.031939 | -0.053817 | 0.0039164
180 | 270 | 90 | -0.36 | 0.74 | 0.36 | -0.74 | -0.2 | 0.45 | 0.15064 | 0.24875 | 0.35064 | -0.20125 |  |  | -0.19559 | -0.086099 | 0.045667
180 | 270 | 270 | 0.47 | 0.22 | -0.47 | -0.22 | 0.63 | -0.07 | 0.27936 | 0.13125 | -0.35064 | 0.20125 |  |  | -0.066866 | -0.20361 | 0.045926
180 | 270 | 300 | 0.52 | 0.35 | -0.52 | -0.35 | 0.68 | 0.06 | 0.47696 | 0.4096 | -0.20304 | 0.3496 |  |  | 0.13073 | 0.074753 | 0.02268
180 | 270 | 330 | 0.35 | 0.36 | -0.35 | -0.36 | 0.51 | 0.07 | 0.50897 | 0.47429 | -0.001035 | 0.40429 |  |  | 0.16274 | 0.13943 | 0.045926
[11 empty] | Coordinates in rotated system |  |  |  | BB coordinates after minimization |  | sum of squars x^2 +y^2 for each table angle
ballCBCT
x (mm) | y (mm) | z (mm)
0 | 0 | 0
CBCT-ball-table0
x | y | z
0 | 0 | 0

## Collimator
[7 empty] | Max square of CA displacement -> |  |  |  | 0.15305 | Solve for smallest max (Collimator Isocentricity)
Collimator |  |  |  |  | CA-X | CA-Z | Coll-X | Coll-Z | CA-X'' | CA-Z'' | CA-R''^2
180 | 0 | 0 | -0.19 | 0.11 | 0.19 | -0.11 | 0.25483 | -0.16986 | -0.064826 | 0.059862 | 0.088237
180 | 90 | 0 | -0.35 | 0.05 | 0.35 | -0.05 |  |  | 0.095174 | 0.11986 | 0.15305
180 | 270 | 0 | -0.16 | 0.29 | 0.16 | -0.29 |  |  | -0.094826 | -0.12014 | 0.15305

## Report
Date and Time | Machine | CBCT-> | CBCT-> | CBCT-> | Table Axis-> | Table Axis-> | Gantry flex | Coll-Gantry-misalign | MLC offset | Table wobble
|  | Gantry Isocenter | Gantry Isocenter | Gantry Isocenter | Gantry Isocenter | Gantry Isocenter |  | Gantry |  | diameter
|  | X (mm) | Y (mm) | Z (mm) | X (mm) | Z (mm) | (mm) | (mm) | (mm) | (mm)
Data Date: 2024-10-15 17:47:43 | TB3 | -0.33 | 0.065 | -0.06625 | 0.016226 | 0.2686 | 0.49 | 0.0575 | -0.2525 | 0.42861

## Instructions
Insert Data:
| 1 | Use Data worksheet
| 2 | use QASRSWL patient; take CBCT and make shifts. After initial CBCT take a second CBCT
| 3 | Shoot MV images in the order shown on the DATA worksheet
| 4 | Find the centroid of the radiation target ball in the contouring spreadsheet.
|  | 4a | Make a high resolution structure
|  | 4b | Set the "apply area" around the ball pretty tight
|  | 4c | Use thresholding contouring; for cermaic ball use threshold of 1000-7000, for tungsten ball use 6000-7000 threshold
|  | 4d | Use centroid report tool in eclipse to find center of ball and origin of acquired image; Make sure to use the correct image, the tool will usually report the data from the first CBCT, and you want data from the last (probably)
| 5 | Find center of ball in second CBCT (after move) and insert ball centroid and origin in A23-C24
| 6 | Copy the data from A1-Y17 area, the order of table, gantry and collimator must be maintained
| 8 | In Analysis worksheet go to the data menu in excel and select solver. If Solver is not installed, install it. Press the solve button.
| 9 | In the Collimator worksheet go to the data menu in excel and select solver. Press the solve button.
| 10 | If everything has gone well the report worksheet should contain the results of the analysis
| 11 | Go to the first sheet, labelled "SNCImport", save it as a CSV, and upload it into SunCHECK