# Linac and Couch Mechanical QA

Last updated by: Cory Knill, Ph.D.

Last update date: September 22nd, 2024

# Overview

- Frequency:

  - Monthly

- Setup/Analysis

  - Placeholder

- Tolerances

- References

  - TG 142: Quality assurance of medical accelerators.

  - TG198: An implementation guide for TG 142.

# Background

Placeholder.

# Procedure

Placeholder

# Tolerances

Placeholder.

# Failure Remediation

Placeholder.

# Calculations

Calculations recorded in SunCHECK are shown in green.

Calculations performed but not recorded in SunCHECK are shown in orange.

Calculations start by taking all the 2D BB-MV offsets from the Winston-Lutz images and converting them to the room coordinate system (where x is crossplane, y is up/down, and z is inplane).

The starting presumption is that the CBCT and subsequent shift has placed the BB at imaging isocenter.

## Gantry

Next, average all the collimator 90 and collimator 270 values at each gantry angle, when the couch is at zero to get the G0, G90, G180, and G270 images.

CBCT – Gantry X = Average of the X BB offsets from G0 and G180 images.

CBCT – Gantry Y = Average of the Y BB offsets from G90 and G270 images

CBCT – Gantry Z = Average of the Z BB offsets from the G0, G90, G180, and G270 images.

Gantry Flex = Difference in z offsets from G0 and G180.

Gantry Isocentricity = Square root of the sum of squares of the maximum – minimum values of X, Y, and Z from the G0, G90, G180, and G270 images.

## Table

The center of rotation of the table in the x and y direction is determined using a built in excel solver, which finds the location that minimizes the maximum Euclidian distance between the center of rotation and any of the bb deviations from the WL images. This is the center of the smallest circle that contains all the BB’s in the couch images.

Table Images = G180C270 – Table at 0, 30, 60, 90, 270, 300, 330 degrees

Table Iso X = X position of smallest circle that contains all the BB’s in the table images

Table Iso Z = Z position of smallest circle that contains all the BB’s in the table images

The table (X,Z) center positions are compared to the gantry (X,Z) isocenter positions and reported.

Table – Gantry Iso X = “CBCT – Gantry X” – “Table Iso X”

Table – Gantry Iso Z = “CBCT – Gantry Z” – “Table Iso Z”

Table Isocentricity = Maximum Euclidean between Table ISO (X,Z) and any BB position on the table images.

## MLC

For all gantry angles take the average X, Y, and Z positions for C90 and 270 images.

X Collimator Position = Average of the X BB offsets from G90 and G270 images at any gantry angle.

Y Collimator Position = Average of the Y BB offsets from G90 and G270 images at any gantry angle.

Z Collimator Position = Average of the Z BB offsets from G90 and G270 images at any gantry angle.

MLC offset = Absolute difference in Z BB position between C90 and C270 images, averaged over all gantry angles.

Coll-Gantry Misalignment = Average of  (X Collimator Position @ G0 – X Collimator Position @ G180) and

(Y Collimator Position @ G90 – Y Collimator Position @ G270)